## Procedure
Converting analog values to electrical conductivity
1) Replace values in sheets 10uScm, 447uScm, and 1413uScm with data collected with your Smart Rock in those solutions
2) From Calibration sheet, use values in cells F4, m, and F5, b, to convert analog 0 values from your Smart Rock data file into electrical conductivity in μS/cm at measured temperature using EC = b*m^analog value

## 10uScm
Packet
Number | Analog[0] | Temperature (°C)
1 | 2821 | 21.72
2 | 2812 | 21.72
3 | 2827 | 21.72
4 | 2819 | 21.72

## 447uScm
Packet
Number | Analog[0] | Temperature (°C)
1 | 7661 | 21.72
2 | 7312 | 21.72
3 | 7711 | 21.72
4 | 7592 | 21.72
5 | 7833 | 21.72

## 1413uScm
Packet
Number | Analog[0] | Temperature (°C)
1 | 9174 | 20.55
2 | 8917 | 20.55
3 | 9100 | 20.64
4 | 9215 | 20.69

## Calibration
Average measured analog values | Average measured temperature (°C) | Solution value at measured temperature (μS/cm) | Solution value at 25°C (μS/cm) | Calibration curve
2819.75 | 21.72 | 8.5642 | 10
7621.8 | 21.72 | 419.732 | 447 | Coefficient m | Coefficient b
9101.5 | 20.6075 | 1297.2779 | 1413 | 1.0008 | 0.89939
Example of applying calibration curve
|  |  |  | Analog value | ECT (μS/cm) | EC25°C (μS/cm)
|  |  |  | 2819.75 | 8.6311 | 9.255
|  |  |  | 7621.8 | 406.0902 | 438.2499
|  |  |  | 9101.5 | 1330.4621 | 1468.8839
〖𝐸𝐶〗_𝑇=𝑏∗𝑚^(𝐴𝑛𝑎𝑙𝑜𝑔 𝑣𝑎𝑙𝑢𝑒)
〖𝐸𝐶〗_𝑇=𝑏∗𝑚^(𝐴𝑛𝑎𝑙𝑜𝑔 𝑣𝑎𝑙𝑢𝑒)
〖𝐸𝐶〗_(25°𝐶)=〖𝐸𝐶〗_𝑇/(1+𝑡𝑐(𝑇−25))

## SolutionTempCompensation
Oakton Conductivity Standard Temperature Compensation |  |  |  |  | Slope | Intercept | adjustment | Temperature coefficient (tc)
Temp °C | 23 uS/cm | 447 uS/cm | 1413 uS/cm | 23 uS/cm | 0.43559 | 12.1032 | -13 | 0.020551
5 | 15.32 | 278 | 896 | 447 uS/cm | 8.8187 | 228.1906 | 0 | 0.022373
10 | 17.11 | 318 | 1020 | 1413 uS/cm | 26.3338 | 754.6048 | 0 | 0.021454
15 | 18.32 | 361 | 1147
16 | 18.65 | 368 | 1173
17 | 18.97 | 376 | 1199
18 | 19.41 | 385 | 1225
19 | 19.85 | 394 | 1251
20 | 20.3 | 402 | 1278
21 | 20.92 | 411 | 1305
22 | 21.51 | 419 | 1332
23 | 22.1 | 430 | 1359
24 | 22.55 | 438 | 1386
25 | 23 | 447 | 1413
26 | 23.43 | 457 | 1440
27 | 23.9 | 467 | 1467
28 | 24.56 | 477 | 1494
29 | 25.2 | 487 | 1521
30 | 25.61 | 496 | 1548
31 | 26.1 | 505 | 1575
Note on J1: Cara Walter: http://www.4oakton.com/Assets/manual_pdfs/35604_00.pdf